## Slide 1: VLF antenna setup manual
University of Oulu
Finland 2021
Claudia Martinez, ISEE Nagoya University

## Slide 2: Receiver Schematics

## Slide 3: Part list
Stay cables x 8 (short x 4, long x 4)
Mast base x 1 (base anchor x 4)
Antenna cables x 2 (NS x 1, EW x 1)
Stay anchor x 4
Mast x 2
Stay connectors x 8
Not pictured: Extra anchors and connector parts

## Slide 4: Set up [1]
Set up the mast base and fix using the 4 anchors
Drive the 4 stay anchors as shown below
North
Orient the base with pivot in the north direction
N
S
E
W
Base
7.5 m
7.5 m
anchors

## Slide 5: Set up [2]
Connect the base and lower mast
Connect the upper and lower masts.
Careful when putting the bolts as too much force can break the mast.
North
bolt and nut
bolt and nut

## Slide 6: Set up [3]
Attach the antenna cables to the top of the mast in their respective directions (follow yellow tape/marker letters)
Make sure the cables don’t cross
S
N
W
E

✗

## Slide 7: Set up [4]
Remove the protection and connect the cable connectors
Attach antenna cables to the mast just above the base (pay attention to directions) being careful with the connectors

## Slide 8: Set up [5]
Attach the short stay cables to the antenna cables
Antenna cable
To anchor
To anchor
To antenna

## Slide 10: Set up [6]
Attach the long stay cables to the middle of the mast from their unadjustable side.
Loosely attach the long stay cables to the anchors on the adjustable side
To the four anchors
To antenna

## Slide 11: Set up [7]
Raise the mast (being careful that the connectors are not in the way)
Check that the loops are not crossed

✗
Start adjusting the length of the stay cables

## Slide 12: Set up [7]
Adjust the cables until the mast is straight and not tilted. Attach antenna cables to anchor points. Loops are extended but not pulling too much on the connectors.

## Slide 13: Set up [8]
Make sure the connectors are covered in waterproof tape and set each connector cable to its mount. Securely fix the connectors using zip ties.

## Slide 14
Final Product

## Slide 15: Set up [9]
Connect each loop antenna to the pre-amplifier making sure they are in their appropriate directions (indicated by tape).
Connect signal cable to amplifier and then to the computer.

## Slide 16: Set up [10]
The gain in the pre-amplifier has been maxed (amp output saturates at +/- 5V) however this might have changed during shipping. Please make sure that the gain is maximized using a screwdriver on the yellow screws moving the screw all the way to 大. If the gain is not maximum, we cannot make measurements of natural VLF waves. When all is ready please put the pre-amp in a weather proof box.
大 (big) 小 (small)
大 bigger
小 smaller

## Slide 17: Note Gain
As the same with attached pictures, please check the "high" gain of the main amplifier in the install.

## Slide 18: PC Set up [11]
Signal cable
to antenna
GPS
Win PC
Main Amp
GPS
antenna
Main Amp
GPS
Channel 1
Channel 2
GPS

## Slide 22
  Conf_VLF.txt
  -nOUJ
  -s02
  -d1F:\vlfdata\
  -a110
  -a210
  -a310
  # -n: Station Name
  # -s: System Number(Antenna Number)
  # -dN: Data spool Directory (N:1-9)
  # -aN: Graph View AmpGain(N:Ch No.1-3)
  Shedule_VLF.txt
  00:00:00,00:00:00
  00:09:50
  00:10:00
  starttime (hh:mm:ss) ,endtime (hh:mm:ss)
  ADtime (hh:mm:ss)
  cycletime (hh:mm:ss)
  To make continuous measurement, set the same time for both starttime and endtime.
  VLF system
  bipolar-5 - +5V Differential Inputs
  40kHz 3ch Sampling
  config files: Conf_VLF.txt and Shedule_VLF.txt at C:\Vb\VLF_A3_40kHz
  data files: F:\vlfdata

## Slide 23: How to restart the system
Turn on the UPS backup battery
Turn on the main amplifier
UPS power backup battery
Push the button for ~3 seconds
until the buzzer sounds,
then turn “INV” LED ON.
switch
switch
backside
Turn on the Windows PC.
Turn on the External Hard Disk Drive.
Automatically turn on when plugged to power cable.
Automatically turn on when PC powered on.

## Slide 24: Data Acquisition Program (VLF_A3_40kHZ.exe)
  Description of the Program
  “C:\Vb\VLF_A3_40kHz\VLF_A3_40kHz.exe” automatically starts up when Windows OS is started.
  When the program is started, it initialize the A/D board, and start data acquisition at the time of starttime.
  The setup parameters of A/D board initialization is as follows:
    Input Range: -5 - +5V
    Memory Type: FIFO
    Clock Type: Internal
    Sampling Clock (usec): 10 (10μsec)
    Start Trigger Type: Software
    Stop Trigger Type: Command
  If you want to start data acquisition manually, choose [Manual] -> [Manual Run] in menu bar. Then the program make data acquisition from the current time over 1 file.
  The data filename is “VLF02_20yymmddhhnnss_40k_Ch.OUJ”
  The acquired data is stored at the external-HDD at F:\ vlfdata\yymmdd\.
  To stop data acquisition, click “stop” button. Then you will see a message like “A/D Stop. 2016/10/15 02:21:01” in the textbox in the leftside. After that, click “Exit”

## Slide 25
  UPS Shutdown Driver
  “Acroware-iGYupsDriver”
  This driver is to shutdown the OS when the electric power stops.
  When the OS is started, this driver is automatically started.
  Service name:AcrowareGupsd
  Startup type: Automatic
  This driver is installed in the following folder
  C:\Program Files (x86)\GSYUASA\AcrowareGYupsd
  Setup Tool
  C:\Program Files (x86)\GSYUASA\AcrowareGYupsd\gyconftool.exe
← COM port number
  Even Log File
  C:\Program Files (x86)\GSYUASA\AcrowareGYupsd\log\EventS.log
  It the connection is correct, the following log is stored in the log file (in Japanese).
  "Acroware UPS シャットダウンドライバ開始(VersionInfo.）"
  "通信に成功しました(COM3)"
  Check of the driver at the time of power stop
  The OS is started to be shutdown 70 second after the power stop.
  The UPS output power will be stoped at 240 second after the start of the OS shutdown.
  When the power is recovered, the UPS output is recovered, and the OS is automatically started.